## Feuille 1
list | causes | gravity | probability | Conclusion | acceptability | Mitigation | probability | acceptability
the patient may fall | the table is turned upside down by the break of one leg | medium (three legs left isostatic) | low | low | A
the patient may fall | the table is turned upside down by the break of two legs | catastrophic | low | Remote | UA | 01.0335 : the legs shall resist 1.2 times the previous value | low | A
| Hitting a wall while entering the X-ray | medium | high | remote | UA | 01.0336 : the system shall have a systematic procedure for entering the X-ray | low | A
the patient might bump his head while in the X-ray | break of the restraints with an unwilling patient | bad | medium | medium | UA | 01.0337 The restraints shall be 1.5 more resistant | low | ALART
Unintended radiation exposure | The disposition of the the table in the X-ray expose more than the legs | catastrophic | very low | low | A
Table unable to enter X-ray | motor fail | catastrophic | low | low